## Slide 1: Impulsivity
Impulsivity: the sudden onset of a sound
  Defined as combination of onset rate (dB/s) and level difference (ΔdB) for periods when gradient exceeds 10 dB/s
Standard: Nordtest Method: NT ACOU 112 (2002)
Hardware requirements
  IEC Class 1 equipment
  Electric background noise 10 dB lower
  than acoustic background noise
Data processing requirements
  A-weighted SPL
  Time weighting “F” or “Fast”
  (125 ms time constant)
  10-25 ms sample period

## Slide 2: Amplitude Modulation
Amplitude modulation: “periodic fluctuations in the level of audible noise from a wind turbine (or wind turbines), the frequency of the fluctuations being related to the blade passing frequency of the turbine rotor(s)”
  Peak to peak amplitude of fundamental modulation frequency and harmonics over a 10 s period, if energy in fundamental frequency is > 4 times the masking level
  AM exists if > 50% of 10 s periods in a 10-minute block meet the above requirement
Standard: Institute of Acoustics Noise Working Group (Wind Turbine Noise) Amplitude Modulation Working Group Final Report: A method for rating amplitude modulation in wind turbine noise (2016)

## Slide 3: Amplitude Modulation
Hardware requirements
  Measurement according to IOA’s Good Practice Guide to the Assessment and Rating of Wind Turbine Noise (IOA NWG, 2013)
  AM quantification method intended to be applied from residential distances, typically greater than 500 m
  lower limit of instrument linearity range shall be no higher than 25 dB(A)
  No specific requirements for wind conditions tested under, but as broad as possible for the site and typical of complaint periods if any
Data processing requirements
  A-weighted 1/3 octave band SPL
  100 ms sample period, 10 consecutive minutes
  3 filter bands considered: 50-200, 100-400, 200-800 Hz

## Slide 4: Low Frequency and Infrasound
Low frequency noise (LFN): 1/3 octave bands centered at 20 – 200 Hz
Infrasound noise: 1/3 octave bands centered < 20 Hz
ILFN: Infrasound + LFN (0.5 – 200 Hz)
Standard: ANSI/ASA S12.9-2016/Part 7: Quantities and Procedures for Description and Measurement of Environmental Sound, Part 7: Measurement of Low-frequency Noise and Infrasound Outdoors and in the Presence of Wind and Indoors in Occupied Spaces

## Slide 5: Low Frequency and Infrasound
Hardware requirements
  Survey grade: general or routine-purpose situations
    LFN frequency response
      +/- 1dB in 1/3 octave bands from 20 – 200 Hz
    ILFN frequency response
      +/- 1dB in 1/3 octave bands from 3.5 – 200 Hz
      +/- 3dB in 1/3 octave bands from 0.5 – 2.5 Hz
    LFN
      The microphone shall include a 175 mm windscreen recommended by the manufacturer or better
    ILFN
      The microphone shall include a 300 mm solid foam (20 ppi) windscreen recommended by the manufacturer or better
    Microphone at a height of 1-1.5 m above ground
    For LFN and ILFN measurements at ground plane and above ground recommended

## Slide 6: Low Frequency and Infrasound
Hardware requirements
  Precision grade: situations requiring the best available precision
    LFN frequency response
      +/- 0.5dB in 1/3 octave bands from 0.4 – 200 Hz
    ILFN frequency response
      +/- 0.5dB in 1/3 octave bands from 0.1 – 200 Hz
    LFN
      The microphone shall include a 300 mm solid foam (20 ppi) windscreen recommended by the manufacturer or better
      The microphone shall be at a height of 1.5 m
    ILFN
      1 m diameter ground plane (like IEC 61400-11) with 90 mm hemispherical wind screen + 450 mm diameter turbulence/secondary wind screen
Data processing requirements
  Calculate background wind noise for each wind speed bin
  Frequency weighting not specified (C, G, or no weighting are more appropriate)